## Slide 1: Syringe Dead Space Volume Table
  Waste volume from dead space should not exceed 0.4 ml total for 6 doses of Pfizer vaccine.
  Example: Syringe/Needle Combination for vial of Pfizer COVID-19 Vaccine
 |  | Needle |  |  |  |  | 
Syringe |  | 0.05 ml |  |  |  |  | 
 |  |  | 0.04 ml | 0.07 ml | 0.07 ml | 0.07 ml | 
 |  |  | 0.06 ml | 0.06 ml | 0.09 ml | 0.08 ml | 
 |  |  | 0.06 ml | 0.09 ml | 0.09 ml | 0.10 ml | 0.12 ml
 |  |  | 0.11 ml | 0.06 ml | 0.11 ml | 0.10 ml | 
 |  |  | 0.13 ml | 0.08 ml | 0.10 ml | 0.11 ml | 
Color | Dead Space
Green | small (≤0.05ml)
Orange | Medium (0.05ml to 0.07ml)
Light Red | Large (0.07ml to 0.10ml)
Dark Red | Xlarge (≥0.11)
Dose | Syringe/Needle (Size) | Dead Space + Uncertainty
1 | Light Red (3 ml) | 0.10 ml + 0.020 ml
2 | Green (1 ml) | 0.05 ml + 0.005 ml
3 | Green (1 ml) | 0.05 ml + 0.005 ml
4 | Green (1 ml) | 0.05 ml + 0.005 ml
5 | Green (1 ml) | 0.05 ml + 0.005 ml
6 | Green (1 ml) | 0.05 ml + 0.005 ml
 | Total Waste | 0.40 ml
Syringe Draw Uncertainty1 | 
Size | Draw Uncertainty
1 ml | +/- 0.005 ml
3 ml | +/- 0.020 ml
Reference:
1) Thobani, S. U., & Steward, D. J. (1992). The accuracy and variability of bolus injections with different sized syringes. Canadian Journal of Anaesthesia, 39(2), 198-201. doi:10.1007/bf03008656
.